## Slide 1: Power measurements vs a-wave
Nick Lalic
VNA Software Developer
Rohde & Schwarz North America

## Slide 2: PA Compression Test
Case study: MACOM Boston, Massachusetts USA
Setup:
  Pulsed RF power
  Driver amplifier
Calibration
  One Path, Two port
  Question: How to power cal?
Input Power (Pin)
Output Power (Pout)
PA S(f, Pin)
Power Cal. Plane
Calibration Planes
Port1
Port2
Driver
VNA

## Slide 3: PA Compression Test
Standard Power cal of source, receiver on Port 1
Calculate compression from S21(f, Pin)
Problems:
  No “2D” source power cal (vs power, frequency)
  Driver amplifier will drift
Option 1 – Power cal
Input Power (Pin)
Output Power (Pout)
S(f, Pin)
Power Cal. Plane
Calibration Planes
Port1
Port2
Driver
VNA

## Slide 4: PA Compression Test
Standard Power cal of source, receiver on Port 1
Enable ALC
Calculate compression from S21(f, Pin)
Problems:
  ALC is slow
  ALC (often) too slow for pulsed RF
Option 2: ALC
Input Power (Pin)
Output Power (Pout)
S(f, Pin)
Power Cal, ALC Plane
Calibration Planes
Port1
Port2
Driver
VNA

## Slide 5: PA Compression Test
Standard Power cal of source, receiver on Port 1
Measure source power (a1) along with S-Parameters
Calculate compression from S21(f, a1)
Option 3: Use measured power
Input Power (Pin)
Output Power (Pout)
S(f, Pin)
Power Cal. Plane
Calibration Planes
Port1
Port2
Driver
VNA

## Slide 6: PA Compression Test
Options 1, 2 illustrated
Measure S21, plot vs theoretical power
Use power cal, ALC to force actual power to theoretical
Option 3: Use measured power
Pin
Gain

## Slide 7: PA Compression Test
Option 3 illustrated
Measure [S] and measured power
Plot vs measured power
Option 3: Use measured power
Pin
Gain

## Slide 8: PA Compression Test
Conclusion: MACOM Boston, MA setup
Input Power (Pin)
Output Power (Pout)
PA S(f, Pin)
S1
a1
b1
S2
a2
b2
Power Cal. Plane
Calibration Planes
Port1
Port2
Driver

## Slide 9: Measurement
VNA cannot plot with measured a-wave x-axis
Software is necessary
Measure:
  Desired parameters (S-Parameter group, S21 etc)
  an(Pn) trace
Plot vs a-wave (measured power)

## Slide 10: Python Example
example.py https://github.com/Terrabits/measure_with_a1
Assumes ‘example1.cal’ contains calibration
Steps:
  Connect to VNA
  Preset
  Setup channel
  Apply cal group ‘example1’
  Setup S21 trace
  Create, display a1 trace
  Sweep
  Query result
  Create plot

## Slide 11: Python Example
Suggested experiment:
  Once with just DUT
  Once with ~ 0.5 dB pad at input to simulate driver amplifier drift

## Slide 12: Contact
Please contact me with feedback or requests
      Nick Lalic VNA Software Developer
      Half Moon Bay, CA
      +1 (424) 200-2846 nick.lalic@rsa.rohde-schwarz.com
      http://vna.rs-us.net